## Slide 1: Metabolomics: Preanalytical Variables
A work in progress

## Slide 2: QTOF Metabolomics
Tandem MS
  Most clinical testing is targeted
  Select by intact mass > fragment > confirm ID by product(s)
  Limited # of analytes at any one time (duty cycle)
QTOF MS
  Time-of-flight (TOF) has mass accuracy to 2ppm
  Compound identified by the accurate intact mass (to 4 decimals)
  No need to fragment*, massive increase in # of analytes detected.

## Slide 3: Targeted and Discovery Metabolomics
Targeted
  138 compounds of interest for inborn errors of metabolism
  Identify by retention time and exact mass
  Quantitate against standards
Discovery
  Extract “features” from QTOF data
  Normalize abundance and align retention times
  Select informative features (relative comparisons)
  Identify compounds based on exact mass
Uncertainty (Quality scores):
Features between samples (RT and mass)
Intensity between conditions (normalization)
Compound identification*

## Slide 4: Study Design
Sample handling affects are significant
  False negatives and false positives from process variability
Plasma from a single individual same day
    Baseline (Fasted, processed and into -80C within 20 min.)
    Delayed Processing (Whole blood at RT 1hr)
    Delayed Freezing (Plasma 4C for 8 hr)
    Standard Freezing (-30C for 14 days)
    Post-prandial (4 hrs after meal)
    Freeze/Thaw (baseline thawed and refrozen)
  3 replicates injected 3 times (9 runs per condition)
    Reversed Phase and HILIC
    Positive ion and negative ion mass spectral acquisition
    216 individual runs

## Slide 5: Targeted Analysis
Standard mix of high calibrators from 4 BGL assays
  Amino acids, acylcarnitines, organic acids, purines/pyrimidines, creatine
  138 compounds in an aqueous matrix
Results
  76 compound tentatively identified
    46 by HILIC Chromatography (+/-)
    35 by Reversed Phase (+/-)
  Issues
    Coelution of isobars (chromatography)
    Multiple peaks for same compound (in source fragmentation)
    Haven’t tried a biological matrix yet
10 min. run time

## Slide 6: Discovery Metabolomics
Q-TOF Data acquisition
  3 replicates of 6 conditions, each injected 3 times
Feature Extraction
  Peak characteristics (shape, height, reproducibility)
  Mass signature (isotopes, adducts)
  Quality score assigned
Filtering and statistical testing on the identified features

## Slide 7: Initial Filter: All samples in at least 1 condition = 1902 features
  B = baseline
  DF = delayed freezing
  DP= delayed processing
  FT = freeze thaw
  P = Post-prandial
  SF = Standard Freezer (-30C)
Log scale
signal intensity
Mean normalized
“Abundance”

## Slide 8: Second Filter: Signal intensity >1,000 (exclude noise) = 1655 features

## Slide 9: Third Filter: CV <85% within condition for all conditions = 1085
Further filters
ANOVA (one way vs Baseline)
  P >0.05
  Corrected for multiple comparisons
  331 features
Fold change from baseline
  >2.0
  182 Features

## Slide 10: Principle Component Analysis
Reduces complex data sets
First 2 to 3 components should explain most of the variance in the data
Helps identify what features are important
Reliability
Magnitude

## Slide 11: Clustering within conditions 182 features statistically different from baseline at p<0.05 and at least a 2 fold difference in abundance
Baseline
Freeze/Thaw
Delayed Processing
Post-prandial
Delayed Freezing
Standard Feezer (-30C)

## Slide 12: Initial Conclusions
We can generate reproducible data (within condition)
Minor effects
  Single freeze/thaw
  1 hr RT prior to spinning
Larger impact
  8hr delayed freezing
  Non-fasting
  -30C storage*
Baseline
Freeze/Thaw
Delayed Processing
Post-prandial
Delayed Freezing
Standard Feezer (-30C)

## Slide 13: Compound Identification
Bottle neck of the process
  Limited by quality of compound databases
  High degree of endogenous modification/degradation
  Only ~30% of compounds with tentative ID
Current data only Reversed Phase Chromatography
  Biased for non-polar lipophilic compounds
  ID’s can be ambiguous, multiple isomers

## Slide 14: Post-Prandial
50 features decreased > 2 fold relative to baseline (and all others)
Only 11 features increased > 2 fold
10/50 reduced compounds are endocannabinoids
Chart: Endocannabinoids
| Baseline | DF | DP | FT | PP | -30C
Anandamide (18:3, n-6) | 76110.05 | 121442.63 | 79380.07 | 84190.734 | 27210.295 | 81867.7
Linoleoyl Ethanolamide | 88424.2 | 91907.805 | 139481.11 | 94099.42 | 33497.215 | 89497.76
Oleoyl Ethanolamide | 387338.22 | 404260.47 | 523679.78 | 398066.97 | 139154.55 | 376273.3
N-(7Z,10Z,13Z,16Z-docosatetraenoyl)-ethanolamide | 122626.2 | 133784.23 | 144969.84 | 151299.38 | 46692.63 | 131655.89
N-Oleoyl-L-Serine
N-arachidonoyl L-serine | 131749.56 | 121054.87 | 137495.16 | 137057.94 | 59327.676 | 130839.13
PGH2-EA | 185589.16 | 177993.55 | 198075.34 | 192968.81 | 84902.53 | 184040.3
Endocannabinoids
  Bind same receptor as THC
  Increase with fasting
  Hunger/satiety signals
  Synthesized “on-demand”
  Short half-life

## Slide 15: Inverse effects of Delayed freezing and -30C storage
44 compounds increased with delayed freezing and decreased with -30C storage
Kamlage et al. 2014 (Clin Chem 60:2)
  11% increased with 16hr 4C plasma storage
  Attributed changes to metabolic activity
Pinto et al. 2014 (Analyst 139)
  Metabolite reductions with 30 day -20C storage
  Attributed to protein precipitation

## Slide 16: Conclusions (so far)
The QTOF MS data is reproducible
Preanalytical variables are critical to success
  Some more than others
Compound identification is a major challenge
The analyses generate a tonne of data
  Garbage in, garbage out
We have a lot still to learn…